Clinical Activities Support Team | East Lancashire Hospitals NHS Trust | Royal Blackburn Hospital | Haslingden Road | Blackburn | BB2 3HH

 Internal: 83377 or 82352

 ExternalLine: 01254 733377

 advancedlifesupportcourses@elht.nhs.uk

The Resuscitation Trolley is an emergency trolley with equipment and drugs to support deteriorating patients with airway issues, fluid management and hypoglycaemia, as well as cardiac arrest. Contents should reflect the checklist. The checklists are official documents and when completed have to be kept for two years to be compliant with the Trust Medical Devices policy.  Anything older than two years can be disposed of.

DAILY CHECKLISTS

- Oxygen contents above half relate to cylinders size D (when full 340 litres) and E (when full 680 litres).  However some areas have large F cylinders in a cradle by the resuscitation trolley.  Wherever possible, it is preferable to use the piped wall oxygen supply.

| OXYGEN CYLINDER | Approximate cylinder life (hrs & min) |
| --- | --- |
| Flow rate litre/minute | 2 L/min | 4 L/min | 6 L/min | 8 L/min | 10L/min | 15L/min |
| Size D ( 340 litre full capacity) | 2¾ hr | 1½ hr | 50 min | 42 min | 34 min | 22 min |
| Size E ( 680 litre full capacity) | 5½ hr | 2¾ hr | 1¾ hr | 1½ hr | 1hr | <½ hr |
| Size F (1360 litre full capacity) | 11¼ hr | 5½ hr | 3¾ hr | 3hr | 2 ¼ hr | 1 ½ hr |

- Tamper tags are obtained from Pharmacy and they must be quick release.   Some areas have been found to be using cable ties which require cutting to gain access to the trolley  –  this is NOT unacceptable.

- 15 litre reservoir mask must available on top of the trolley.

WEEKLY CHECKLISTS

First Drawer

- Replacement laryngoscope and blade is now an integrated unit. It is acceptable to continue to use in-date individual disposable laryngoscope handle and blades if you have them.

- Suction catheters must be compatible with your suction unit. Commonly wards stock the Aeroflow suction catheters with the male end connection and not the blunt end catheters which are not compatible to the suction tubing without an adaptor

- The Vygon FG16 grey and FG18 green cannulae are in the airway drawer and available for thoracocentesis decompression.

- Other cannulae for IV therapy are in the second drawer.

Second Drawer

IV cannulae & flushes

Third Drawer

The third drawer is identified for drugs and fluid.

The blue sealed drug box must be stored in this drawer.  Adrenaline 1:1000 for anaphylaxis is in the blue box but Secondary drugs for anaphylaxis   Chlorphenamine 10mg injection &  Hydrocortisone 100mg injection should be in the drug cupboard.

The Hypoglycaemia box & Hypoglycaemia algorithm should be in the 3rd drawer

Additional fluid and drugs, not identified on the checklist, should not be on the trolley.

Replacement Resuscitation trolley equipment

Royal Blackburn Teaching Hospital

Replacement equipment for the top of the resuscitation trolley or the Airway drawer can continue to be obtained from AMU(A) 24/7.  Any equipment taken to replenish the resuscitation trolley must be signed for in the green folder.

Burnley General Teaching Hospital

All BGH ward and clinic areas will be able to obtain replacement equipment for the top of the resuscitation trolley or the Airway drawer from Obstetric and Gynaecology theatre.

Community Hospitals

Pendle Community Hospital, Clitheroe Community Hospital and Accrington Victoria Hospital can continue to obtain relevant Airway equipment from AMU(A) or contacting the Clinical Activities Support Team.  CAST cannot assist in providing any ‘regular’ equipment that the ward can order independently   e.g. Aeroflow suction catheters.

If you have any queries related to the provision of Airway replacement stock please contact advancedlifesupportcourses@elht.nhs.uk